Moscow State University

Sternberg Astronomical Institute

Specification

2.5 m telescope with enclosure

version 7.6

| Number | Item description, type | Quantity, notes |
| --- | --- | --- |
| 1000 | 2.5 m mirror telescope system | 1 |
|  | optics and related mechanical characteristics: |  |
| 1001 | Primary mirror M1, with light diameter 2.5m, fabricated from the material with near zero coefficient of expansion (Sitall or Zerodur). The mirror cell (the support, ventilation, or cooling system) and the shape (lightweighting) must provide the short time of relaxation to environment conditions to avoid mirror seeing. |  |
| 1002 | Convex secondary mirror M2 fabricated from similar material to the primary provides telescope focal ratio F/8. Central obscuration of no more than 15% area is allowed including blinds. The secondary mirror cell is mounted on a remotely controlled hexapod. |  |
| 1003 | Two mirror system must be optimized without a corrector to have the flat field of view of 10 arcmin diameter and with the image quality better than 80% encircled energy in 0.2 arc seconds diameter. |  |
| 1004 | The support systems of the primary and secondary mirrors must provide the demanded image quality non depending on the telescope elevation. |  |
| 1005 | One Cassegrain focal station C1 and four Nasmyth focal stations (ports) must be provided. Two main Nasmyth stations are located at the central holes of the elevation axles and two additional (“student ports”) are oriented at 90 and 270 degrees with respect to the elevation axis. First main Nasmyth station N1 must provide the unvignetted field of view (FOV) 40 arcmin in diameter; opposite station N2 must provide FOV 15 to 20 arcmin (depending on axle mechanical restrictions), the additional ports N3 and N4 all have 15 arcmin FOV. The optical clearance for the C1 station must provide at least 1 degree unvignetted FOV. | N3 and N4 ports are optional |
| 1006 | Nasmyth focal stations must be provided with a deployable tertiary mirror M3 assembly allowing switching between Cassegrain and any of the Nasmyth focal station during normal telescope operation. M3 is fabricated from the material similar to the primary. Its assembly must not give additional light obscuration or vignetting when passing light to Cassegrain station. Switching should take not more than 2 minutes and must provide the pointing repeatability not worse than 2 arcsec rms (max deviation 6 arcsec) for any Nasmyth station. |  |
| 1007 | All telescope mirrors must be coated high reflecting Al (R>88%) with a protecting SiO2 layer. Work range of the telescope is optical and near-infrared (0.3 – 2.2 mu). |  |
| 1008 | A wide field corrector assembly at the Cassegrain focal station C1 must provide an unvignetted corrected FOV 40 arc minutes diameter with optimal image quality (better than 80% encircled energy in 0.3 arc seconds diameter). The corrector must be fabricated from fused quartz and optical surfaces must be AR coated optimized for blue spectral region. Corrector must be engineer-mode (day-time) installable/removable. |  |
| 1009 | The wide field corrector similar to that for Cassegrain station must fit the wide-field Nasmyth station N1 and must be engineer-mode removable. |  |
| 1010 | The removable automatic Atmospheric Dispersion Corrector with FOV not less than 12 arcmin is provided in station N2 for needs of long-slit spectroscopy. | option |
|  | Instrument stations and interfaces: |  |
| 1020 | The Cassegrain focal station is equipped with an instrument adaptor containing the mechanical field derotator, rotating instrument interface (flange), and acquisition/guidance unit (AGU). Field derotator must provide the unguided rotation tracking accuracy 0.15 arcsec rms at a distance 20 arcmin from the center of FOV (0.4 arcmin of position angle) during 1 hour integration at any allowed azimuth and elevation. AGU must be capable of selecting the guide star in the distance range from 5 arcmin from center of the FOV to its boundary. |  |
| 1021 | The wide-field Nasmyth station N1 is equipped with the instrument adaptor (derotator, instrument interface, AGU) similar or identical to the adaptor of Cassegrain station. |  |
| 1022 | The narrow-field Nasmyth station N2 is equipped with the optical field derotator compatible with the ADC unit. It must be engineer-mode removable |  |
| 1023 | Additional Nasmyth stations N3 and N4 represent openings in the central section of the tube with 4 instrument fixing holes (thread M10) and closed with lids when not in use. |  |
| 1024 | Wide field corrector(s) (when installed) must not extend beyond the instrument flange. Backfocus distance is 300 to 500 mm in C1 and not less than 200 mm in N1, N2. Maximal load for C1 station is 250 kg, for N1, N2 stations is 400 kg and for N3, N4 is 25 kg. C1 clearance is not less than 1.5 m; N1, N2 must provide space 1 m from instrument flange along light beam and +/- 75 cm across beam horizontal and 1 m down to the platform. |  |
|  | Tube |  |
| 1030 | The tube is assumed made of Serrurier trusses. Mirrors M1 and M3 must have remotely driven covers. The M2 and near-M1 baffles should be designed to prevent the stray light in detector area in Nasmyth and Cassegrain stations. |  |
| 1031 | The M2 unit must provide the fine focussing with a provision of the automatic focus correction using the temperature data from the distributed sensors and/or from the AGU image data. |  |
| 1032 | M1, M2 and M3 cells must allow easy removal/installation of mirrors for reflective layer re-coating. Corresponding means (damped protection cases, covers etc.) for their safe transportation to another observatory coating plant must be provided. |  |
|  | Mount |  |
| 1100 | The telescope mount is alt-azimuth. The azimuth axis will have hydrostatic bearings and will provide the azimuth range -270 degrees to +270 degrees. The elevation axis on roller (or hydrostatic) bearings must provide smooth non-sticking tracking and least problematic maintenance; the elevation range must be 0 to +90 degrees. The first eigenfrequency of at least 7Hz must be provided by the rigid mount structure. |  |
| 1101 | The zenith blind spot will be no larger than 2 degrees around zenith. |  |
| 1102 | The pointing of at least 3 degree per second must be provided. The pointing accuracy of the telescope must be better than 3 arcsec rms (max deviation 5 arcsec). |  |
| 1103 | Unguided tracking accuracy over a period of 600 seconds must be 0.2 arc seconds rms (max deviation 0.4 arcsec); guided tracking accuracy over a period of 3 hours must be 0.2 arc seconds rms. |  |
|  | Telescope control system and environment conditions |  |
| 1201 | The control system hardware (power electronic units) must be located in racks which are assembled in a separate control room (some 10 meters from the dome). All the heat sources which cannot be detached from the mount must be carefully isolated and/or cooled to prevent heat fluxes from surfaces. |  |
| 1202 | Cable channels in the mount must provide the space for the control lines of the attached instrumentation (optical fibre lines, twisted pairs etc.) from the instrument stations to the control room. |  |
| 1203 | The specialized computers must be provided for the telescope control needs having the high reliability level (industrial standards) which will be located in the control room. |  |
| 1204 | It is assumed that the telescope must operate to specification within the operating conditions: -20 degrees centigrade to +25 degrees centigrade; relative humidity 5% to 90%; wind speed to 60 km/hour; maximal snow layer not more than 1 meter. The telescope installation is supposed near the Kislovodsk Solar Station, 25 km from Kislovodsk, at an altitude of 2100m (N 43o44', E 42o 32'). |  |
| 1300 | Telescope spares and consumables for two years, tooling for maintenance | 1 |
| 1400 | Technical service contract for the telescope for an initial five year period with yearly visits. | 1 |
| 1401 | Education of the russian personnel (3-5 persons) during 2 weeks on a telescope assembly stage. |  |
|  | Instrumentation |  |
| 2000 | Scientific CCD camera 4Kx4K with shutter and filters | 1 |
| 2001 | Based on 2 back-illuminated buttable 15mu pixel devices (e.g. CCD44-82bi) with astronomy broad band AR coating. Readout noise not more than 3e. |  |
| 2002 | The camera controller electronics delivers a digitisation depth of 16 bits, programmable readout rate, providing readout noise less than 3e at 200 Ks/s, programmable camera gain. Cable between camera and PC must be not less than 50 meter long. |  |
| 2003 | High-reliability low-maintenance cryostat with closed cycle cooler of CRYOTIGER type (cools down to 140K). |  |
| 2004 | Two computers under Linux OS must provide data acquisition and data storage in FITS format. |  |
| 2005 | Camera shutter must provide unvignetted CCD field, minimal exposure 100 ms and equality of exposures in any field point with 2 ms precision. |  |
| 2006 | A filter wheel assembly mechanism must contain a double wheel, each capable of independent operation, and each with 5 (or 6) positions, capable of accommodating filters 100 mm in diameter. Set of U,B,V,R, I, H-alpha (fwhm 5nm @ 656 nm) filters must be supplied. Filters are operated remotely. Change between adjacent filters must take 1-3 seconds. Filter setting precision must provide repeatability of not worse than 10 mu. |  |
| 2100 | Long-slit spectrograph | 1 |
| 2101 | A provision of a long-slit low-resolution high-throughput spectrograph is to be considered. Spectrograph must provide several resolutions of 1500 to 10000 with help of a holographic grating dispersion element. Collimator is a parabolic f/8 mirror, camera lens must provide the R=10000 resolution in spectral region 300 – 11000 nm on backside illuminated CCD 2K x 4K (e.g. CCD44-82bi) cooled by CRYOTIGER. Additional CCD slit-viewing camera must provide guiding and centring the target. Several slits of different width and with a beam splitter/condenser option are installed on a motorized wheel and are inclined to let the reflected light in a CCD camera via a suitable re-imaging lens. The pre-slit unit must reserve the beam escaping (using insertable 45 degree mirror) into the future high resolution spectrograph installed nearby and contains also the calibration lamps set. Detailed specifications will follow later. | option |
| 2200 | Near-infrared camera | 1 |
| 2201 | Camera 1K x 1K with a J,H,K-filters wheel (1.2,1.6,2.2 mu) with a cryostat and controller to be installed in C1 port. | option |
| 2300 | Observations support automatic telescope with cameras | 1 |
| 2301 | Reflecting telescope on an automatic fast equatorial mount (pointing 10 degree/sec with 10 arcsec rms precision) with D=500mm and f/3.8. Will be used for photometric support of 2.5m observations (extinction monitoring, synchronized multi-channel photometry etc.) and installed outside its enclosure. | option |
| 2302 | Peltier cooled front-illuminated CCD cameras 4Kx4K Alta16U grade 2. | option 4 |
| 3000 | Enclosure | 1 |
| 3001 | The dome of an appropriate size must be provided. The fixed part maintains the suitable support of the rotating part and will be installed on a concrete walls of the observatory building designed according to the telescope and dome specifications. Rotation with speeds 0.01 to 4 deg/sec; slit open/close time 2 minutes or less. |  |
| 3002 | Dome walls are clad with 10cm foam-like insulation to lower the power consumption of the climatic system of the dome interior. External surface has a high reflection and durability coating. Dome and telescope structure and power systems must be optimized for preventing the significant dome seeing with help of the remote controllable louvers and/or ventilation system. The ventilation system must provide continuous air extracting from the dome during observations. Climate-control system must provide the interior temperature stabilized with the next night expected temperature. |  |
| 3003 | The crane for mirrors removal/installation and telescope maintenance must be provided, as well as the carriages and tooling for mirrors and instruments transportation. All the technological process of the mirrors (de)installation for the surfaces recoating must be supported in details. |  |
| 3004 | Dome must withstand wind gusts of up to 30m/s when open and 50m/s when closed. Rain-proof for storms up to 20m/s. |  |
| 3005 | The maximal detected earthquake load during last century in the 100km neighborhood is M=6.5 magnitude (Richter scale); the Maximal likelihood earthquake (probability of exceeding during 50yr is 10%) is about M=7 with the peak ground acceleration of 3.2-4m/s^2. |  |
| 3006 | Mechanical interfaces of the enclosure and telescope with the dome building must be provided. The specifications, drawings and all the relevant information must be provided during first 6 months of the contract to enable the development of the construction documentation for the dome building and concrete telescope pier. |  |
|  | Auxiliary components |  |
| 4000 | Hardware for remote and fully robotic operation of the telescope. Engineering panel for driving the telescope and enclosure. | 1 |
| 4010 | Reference time system providing high accuracy UT (better than 0.1 ms, e.g. based on a GPS receiver) | 1 |
| 4020 | Automatic meteostation must be provided and installed outside the enclosure at a 20m distance from telescope and 5m elevation and connected to the Environment Control System. | 1 |
| 4030 | Complete cabling set for the telescope and dome control and scientific equipment | 1 |
|  | Software |  |
| 5000 | Telescope Control System (TCS) software and firmware. Realtime part must provide correct, fast and robust management of axes encoders and motors and other less response-demanding system components. Safety system must provide protection against collisions and proper interlocks. Upper level user interface software must work under OS Linux. | 1 |
| 5010 | Dome Control System (DCS) for managing the rotation (synchronized with telescope), slit opening/closing, ventilation, air conditioning (maintaining the thermal equilibrium of the telescope mount and mirrors with the next observing night expected temperature) and mirror covers. | 1 |
| 5020 | Autoguider Control System (ACS) for selection of a guiding star near the science field, driving the AGU sliding mechanism to it and guiding during all the scientific observations | 1 |
| 5030 | Instrument Control System (ICS) prototypes for the supplied instrumentation (imaging camera with a shutter and filters assembly and others) as examples of the integration of the instrument software in the single system driven by OCS. | 1 |
| 5040 | Meteostation and weather forecast interface for handling the observing conditions (ECS – environment control system) | 1 |
| 5050 | Observatory Control System (OCS) which is the highest level software integrating the TCS, DCS, ACS, ECS and ICS components which allows the fully remote (via Internet) controllable system functioning and provides the user with convenient interface for observations management. | 1 |
| 5099 | The upper logical level software must be provided with open sources. |  |
| 6000 | Full site installation of the 2.5 m telescope and enclosure. | 1 |
| 6001 | Transport to Kislovodsk (railroad) and further to the Kislovodsk Solar Station, 25 km from Kislovodsk, by auto road at an altitude of 2100m, last 5 km are unpaved. |  |
| 6002 | Final alignment, testing, commissioning |  |
|  | Documentation |  |
| 7000 | Complete technical documentation with drawings for telescope, enclosure and related hard- and software; the maintenance manual and schedule. |  |